## Attenuation Varying Input Power
|  | Frequency (Hz) | Noise Attenuation Factor
|  |  | 0.5 mW Input Power | 1.0 mW Input Power | 1.5 mW Input Power | 3 mW Input Power | 5 mW Input Power | 10 mW Input Power | 30 mW Input Power | 100 mW Input Power
|  | 6 | 181.4097 | 155.964 | 86.9545 | 134.1944 | 98.5177 | 334.8938 | 295.5279 | 451.5301
|  | 8 | 194.499 | 252.5362 | 88.7948 | 147.6788 | 96.0868 | 331.7624 | 243.7546 | 508.9219
|  | 10 | 225.1182 | 258.0258 | 78.926 | 165.3497 | 100.4862 | 297.4656 | 226.2063 | 417.8221
|  | 20 | 139.3982 | 238.3023 | 138.511 | 183.075 | 124.5794 | 290.7342 | 213.5054 | 241.346
Noise Eater Noise Attenuation Factor as a Function of Input Power |  | 40 | 82.9357 | 172.2283 | 92.9918 | 166.6119 | 131.6415 | 402.7018 | 209.1597 | 129.8864
|  | 60 | 54.9051 | 137.1426 | 89.3523 | 137.7104 | 133.0616 | 231.3093 | 101.7204 | 86.9758
|  | 80 | 42.2673 | 99.1585 | 84.9964 | 103.8346 | 94.6561 | 168.3996 | 77.9543 | 67.9824
Item # | NEL03A(/M) | 100 | 33.423 | 81.4812 | 82.1766 | 83.8873 | 113.4408 | 159.642 | 61.1546 | 54.4574
DISCLAIMER: The data presented here are typical. Slight variations in performance data will occur from lot to lot. Please contact Technical Support with any questions regarding the use or reliability of this data. |  | 120 | 28.1377 | 91.7931 | 76.3126 | 78.8235 | 90.9568 | 118.8063 | 53.8292 | 46.3162
|  | 140 | 24.1669 | 57.4 | 73.4345 | 80.4899 | 93.5747 | 100.0577 | 46.5829 | 40.6715
|  | 160 | 21.2048 | 50.9919 | 71.7619 | 69.1897 | 85.9226 | 88.8187 | 39.7063 | 35.8776
|  | 180 | 18.8811 | 45.4896 | 70.5299 | 58.4592 | 82.5186 | 80.6147 | 35.831 | 31.0163
|  | 200 | 17.1212 | 40.7306 | 60.2398 | 55.6674 | 78.8287 | 73.5918 | 32.0072 | 28.0279
|  | 200 | 16.9855 | 40.361 | 61.8831 | 53.156 | 75.7555 | 70.3835 | 32.2732 | 28.149
|  | 300 | 11.6538 | 28.4178 | 43.743 | 37.7745 | 66.2357 | 47.8783 | 21.5658 | 18.7823
Additional Information: |  | 400 | 8.9148 | 21.3966 | 33.464 | 30.4138 | 60.53 | 35.3898 | 15.9308 | 13.859
Data taken at 785 nm, 5% signal modulation (noise level) |  | 500 | 7.061 | 16.7858 | 26.0667 | 24.4444 | 52.9745 | 27.3601 | 12.359 | 10.7397
|  | 600 | 5.8155 | 13.8032 | 21.4438 | 20.5345 | 46.6624 | 21.7692 | 9.9626 | 8.7065
|  | 700 | 4.9194 | 11.6072 | 18.02 | 18.1443 | 42.0711 | 17.3324 | 8.1663 | 7.2351
|  | 800 | 4.1729 | 9.8773 | 15.4259 | 15.3584 | 37.536 | 12.2692 | 6.2851 | 6.1057
|  | 900 | 3.5902 | 8.5359 | 13.3666 | 13.3742 | 33.5602 | 8.481 | 4.4792 | 5.2219
|  | 1000 | 3.1292 | 7.4137 | 11.6555 | 12.1607 | 29.9888 | 6.1653 | 3.4531 | 4.4966
|  | 1200 | 2.4236 | 5.6277 | 8.9195 | 9.2276 | 23.6896 | 4.0199 | 2.5896 | 3.3646
|  | 1400 | 1.9449 | 4.3479 | 6.8389 | 7.204 | 18.1243 | 2.5362 | 1.6122 | 2.5669
|  | 1600 | 1.6116 | 3.3174 | 4.9508 | 5.2973 | 13.8768 | 1.9676 | 1.4072 | 2.0071
|  | 1800 | 1.3712 | 2.613 | 4.3877 | 3.466 | 13.3315 | 1.7108 | 1.2886 | 1.7486
|  | 2000 | 1.0771 | 2.384 | 2.622 | 3.4159 | 7.6036 | 1.7826 | 0.92461 | 1.3244
|  | 2200 | 0.92742 | 2.2221 | 2.502 | 1.9405 | 5.2389 | 1.6308 | 0.79445 | 1.0372
|  | 2400 | 0.84406 | 1.5991 | 2.7488 | 1.5843 | 4.6718 | 1.1645 | 0.7423 | 0.70819
|  | 2600 | 0.81178 | 1.1475 | 2.6096 | 1.7125 | 4.3746 | 1.0038 | 0.72139 | 0.57338
|  | 2800 | 0.78582 | 0.96519 | 1.3328 | 1.9251 | 3.7635 | 0.93689 | 0.70958 | 0.58126
|  | 3000 | 0.76875 | 0.90526 | 1.0661 | 1.6018 | 2.4885 | 0.89841 | 0.70799 | 0.53378
|  | 3200 | 0.76139 | 0.85131 | 0.96155 | 0.83556 | 2.0414 | 0.87142 | 0.70483 | 0.50868
|  | 3400 | 0.75694 | 0.82403 | 0.89485 | 0.6987 | 2.1035 | 0.85462 | 0.70348 | 0.48821
|  | 3600 | 0.75798 | 0.81039 | 0.86012 | 0.68334 | 1.9792 | 0.85686 | 0.70882 | 0.53212
|  | 3800 | 0.75677 | 0.7993 | 0.84153 | 0.62606 | 2.153 | 0.83553 | 0.70585 | 0.4638
|  | 4000 | 0.75881 | 0.79706 | 0.83425 | 0.61293 | 1.9721 | 0.83406 | 0.70806 | 0.44942

## Attenuation Varying Input Mod.
|  | Frequency (Hz) | Noise Attenuation Factor
|  |  | 5% Modulation | 10% Modulation | 15% Modulation | 20% Modulation | 30% Modulation
|  | 6 | 134.1944 | 154.3011 | 125.7225 | 122.4001 | 161.0803
|  | 8 | 147.6788 | 179.8118 | 136.4868 | 122.0365 | 180.0936
|  | 10 | 165.3497 | 253.6389 | 134.5269 | 126.1351 | 166.7114
|  | 20 | 183.075 | 180.422 | 127.2258 | 140.0338 | 153.5918
Noise Eater Noise Attenuation Factor as a Function of Input Signal Modulation (Noise Level) |  | 40 | 166.6119 | 232.1617 | 129.0099 | 121.8885 | 126.3459
|  | 60 | 137.7104 | 128.1411 | 109.5663 | 107.1065 | 112.1212
|  | 80 | 103.8346 | 129.2412 | 92.6344 | 93.0921 | 93.7182
Item # | NEL03A(/M) | 100 | 83.8873 | 99.4013 | 89.7264 | 81.4085 | 83.7529
DISCLAIMER: The data presented here are typical. Slight variations in performance data will occur from lot to lot. Please contact Technical Support with any questions regarding the use or reliability of this data. |  | 120 | 78.8235 | 84.1858 | 76.3416 | 74.578 | 71.3901
|  | 140 | 80.4899 | 79.351 | 71.2667 | 67.9884 | 63.2451
|  | 160 | 69.1897 | 70.9614 | 65.2716 | 62.01 | 55.6494
|  | 180 | 58.4592 | 64.0885 | 59.1387 | 57.49 | 48.3111
|  | 200 | 55.6674 | 59.1046 | 55.4913 | 52.9321 | 37.3514
|  | 200 | 53.156 | 59.7459 | 55.0366 | 52.3697 | 37.5386
|  | 300 | 37.7745 | 42.3526 | 40.597 | 37.5483 | 6.182
Additional Information: |  | 400 | 30.4138 | 34.1687 | 32.0453 | 21.3257 | 3.4372
Data taken at 785 mn, 3 mW input power. |  | 500 | 24.4444 | 28.1809 | 25.1885 | 7.9032 | 2.5894
|  | 600 | 20.5345 | 23.841 | 15.2896 | 4.8189 | 2.1374
|  | 700 | 18.1443 | 20.1047 | 8.2855 | 3.5434 | 1.8511
|  | 800 | 15.3584 | 16.547 | 5.3234 | 2.814 | 1.66
|  | 900 | 13.3742 | 12.8127 | 3.8683 | 2.3503 | 1.5285
|  | 1000 | 12.1607 | 8.9172 | 3.0966 | 2.1097 | 1.4088
|  | 1200 | 9.2276 | 4.6148 | 2.2486 | 1.7447 | 1.1302
|  | 1400 | 7.204 | 3.145 | 1.9793 | 1.5182 | 1.0094
|  | 1600 | 5.2973 | 2.5409 | 1.6524 | 1.162 | 0.95211
|  | 1800 | 3.466 | 2.3839 | 1.3939 | 1.0505 | 0.92887
|  | 2000 | 3.4159 | 1.8423 | 1.0584 | 0.96235 | 0.89143
|  | 2200 | 1.9405 | 1.5141 | 0.95331 | 0.91859 | 0.87382
|  | 2400 | 1.5843 | 1.0935 | 0.90313 | 0.89521 | 0.86318
|  | 2600 | 1.7125 | 0.93733 | 0.86742 | 0.87649 | 0.85464
|  | 2800 | 1.9251 | 0.86316 | 0.84473 | 0.86315 | 0.84904
|  | 3000 | 1.6018 | 0.82479 | 0.83561 | 0.85606 | 0.84515
|  | 3200 | 0.83556 | 0.79688 | 0.81949 | 0.84911 | 0.84167
|  | 3400 | 0.6987 | 0.77765 | 0.81116 | 0.84409 | 0.83948
|  | 3600 | 0.68334 | 0.79733 | 0.81718 | 0.84383 | 0.83959
|  | 3800 | 0.62606 | 0.75779 | 0.80263 | 0.83819 | 0.83682
|  | 4000 | 0.61293 | 0.74685 | 0.79978 | 0.83586 | 0.83598

## Attenuation Varying Output Powe
|  | Frequency (Hz) | Noise Attenuation Factor
|  |  | 10% Signal Attenuation | 20% Signal Attenuation | 30% Signal Attenuation | 40% Signal Attenuation | 50% Signal Attenuation | 60% Signal Attenuation | 70% Signal Attenuation
|  | 6 | 134.1944 | 137.8149 | 142.6038 | 94.3784 | 95.3312 | 146.2313 | 151.1373
|  | 8 | 147.6788 | 168.4674 | 119.7867 | 116.1861 | 127.0797 | 114.8901 | 167.9058
|  | 10 | 165.3497 | 148.8071 | 126.8302 | 95.3557 | 123.1655 | 128.2818 | 176.353
|  | 20 | 183.075 | 141.6006 | 136.9537 | 120.5746 | 132.2632 | 167.353 | 136.005
Noise Eater Noise Attenuation Factor as a Function of Output Power Level (Attenuation Factor) |  | 40 | 166.6119 | 140.1115 | 161.6585 | 110.7609 | 128.8444 | 144.6729 | 102.2249
|  | 60 | 137.7104 | 140.1085 | 184.8244 | 103.4543 | 154.9058 | 97.0738 | 130.8507
|  | 80 | 103.8346 | 111.0541 | 88.411 | 84.2556 | 95.0489 | 88.0388 | 83.3371
Item # | NEL03A(/M) | 100 | 83.8873 | 85.7674 | 93.3228 | 75.8138 | 75.3315 | 89.5867 | 53.0747
DISCLAIMER: The data presented here are typical. Slight variations in performance data will occur from lot to lot. Please contact Technical Support with any questions regarding the use or reliability of this data. |  | 120 | 78.8235 | 82.9082 | 78.7684 | 68.6294 | 69.1001 | 62.6435 | 58.4039
|  | 140 | 80.4899 | 76.8376 | 73.7653 | 62.3376 | 70.1743 | 55.0032 | 61.2113
|  | 160 | 69.1897 | 70.398 | 66.45 | 63.4298 | 65.2212 | 56.5958 | 64.3516
|  | 180 | 58.4592 | 62.5866 | 59.8214 | 59.9391 | 56.3377 | 49.8142 | 41.2674
|  | 200 | 55.6674 | 57.1488 | 54.4979 | 51.8758 | 46.1204 | 43.2394 | 33.6971
|  | 200 | 53.156 | 53.4626 | 53.937 | 51.7643 | 49.3264 | 41.4286 | 37.2274
|  | 300 | 37.7745 | 38.1467 | 38.0483 | 37.0791 | 32.4924 | 29.6723 | 25.6567
Additional Information: |  | 400 | 30.4138 | 31.1606 | 29.8443 | 29.6039 | 26.2731 | 23.2103 | 18.9311
Data taken at 785 nm, 3 mW input power, and 5% signal modulation (noise level) |  | 500 | 24.4444 | 24.8223 | 24.6946 | 23.8634 | 21.002 | 17.8929 | 14.8286
|  | 600 | 20.5345 | 20.7431 | 20.2767 | 20.0322 | 17.1852 | 15.1172 | 12.2346
|  | 700 | 18.1443 | 17.9189 | 17.4301 | 17.3591 | 15.2747 | 13.0176 | 10.8824
|  | 800 | 15.3584 | 15.07 | 14.8232 | 14.5683 | 12.7844 | 11.2083 | 9.0802
|  | 900 | 13.3742 | 13.1702 | 12.9459 | 12.6032 | 11.1606 | 9.6867 | 7.8795
|  | 1000 | 12.1607 | 11.6177 | 11.322 | 11.0813 | 9.7658 | 8.5836 | 6.9326
|  | 1200 | 9.2276 | 9.0538 | 8.8366 | 8.6972 | 7.7323 | 6.7903 | 5.4183
|  | 1400 | 7.204 | 6.4434 | 6.1855 | 6.5344 | 6.1944 | 5.5265 | 4.2279
|  | 1600 | 5.2973 | 5.0947 | 5.0436 | 5.1575 | 4.6174 | 4.1744 | 3.1347
|  | 1800 | 3.466 | 3.7558 | 4.0856 | 7.604 | 5.6551 | 3.2236 | 2.6202
|  | 2000 | 3.4159 | 3.3032 | 3.2589 | 3.303 | 3.2645 | 2.2854 | 1.8565
|  | 2200 | 1.9405 | 1.9583 | 1.9071 | 2.0173 | 2.0484 | 1.8353 | 1.7601
|  | 2400 | 1.5843 | 1.5791 | 1.6189 | 1.7003 | 1.7687 | 1.775 | 1.712
|  | 2600 | 1.7125 | 1.7289 | 1.7819 | 1.8394 | 1.8878 | 1.7772 | 1.4217
|  | 2800 | 1.9251 | 1.929 | 1.9202 | 1.8677 | 1.8903 | 1.3193 | 0.97161
|  | 3000 | 1.6018 | 1.4182 | 1.3362 | 1.2711 | 1.1919 | 0.9145 | 0.73421
|  | 3200 | 0.83556 | 0.80247 | 0.78426 | 0.75828 | 0.7342 | 0.65024 | 0.5944
|  | 3400 | 0.6987 | 0.67362 | 0.66948 | 0.64767 | 0.63959 | 0.58664 | 0.55351
|  | 3600 | 0.68334 | 0.67476 | 0.70586 | 0.68783 | 0.65865 | 0.5875 | 0.60495
|  | 3800 | 0.62606 | 0.61781 | 0.62146 | 0.6153 | 0.61201 | 0.57508 | 0.56589
|  | 4000 | 0.61293 | 0.61307 | 0.62645 | 0.61582 | 0.61858 | 0.58532 | 0.56376

## Modulation Performance
|  | Frequency (Hz) | Max Modulation Depth
|  | 10 | 100
|  | 15 | 98.9119
|  | 20 | 97.2726
|  | 25 | 95.309
Noise Eater Modulation Performance |  | 30 | 92.5425
|  | 35 | 89.6161
|  | 40 | 83.2929
Item # | NEL03A(/M) | 45 | 77.1224
DISCLAIMER: The data presented here are typical. Slight variations in performance data will occur from lot to lot. Please contact Technical Support with any questions regarding the use or reliability of this data. |  | 50 | 71.9031
|  | 55 | 66.9041
|  | 60 | 62.5701
|  | 65 | 58.6159
|  | 70 | 55.1742
|  | 75 | 51.775
|  | 80 | 49.0036
Additional Information: |  | 85 | 45.9777
|  | 90 | 43.5385
|  | 95 | 41.4
|  | 100 | 39.3794
|  | 100 | 39.4646
|  | 150 | 25.8836
|  | 200 | 18.9328
|  | 250 | 14.8519
|  | 300 | 12.3198
|  | 350 | 10.7504
|  | 400 | 9.5448
|  | 450 | 8.6147
|  | 500 | 7.9089
|  | 500 | 7.8894
|  | 600 | 6.804
|  | 700 | 6.0352
|  | 800 | 5.429
|  | 900 | 4.978
|  | 1000 | 4.5559
|  | 1000 | 4.5442
|  | 1500 | 3.2938
|  | 2000 | 2.6148
|  | 2500 | 2.2385
|  | 3000 | 1.9436

## Long-Term Performance
|  | Time (Hours) | Normalized Signal
|  |  | Without Noise Eater | With Noise Eater
|  | 0 | 0.99012 | 0.8091
|  | 0.02889 | 0.98077 | 0.80882
|  | 0.0575 | 0.9787 | 0.80859
|  | 0.08611 | 0.9906 | 0.809
Noise Eater Noise Long-Term Performance |  | 0.11472 | 0.98488 | 0.80873
|  | 0.14389 | 0.97816 | 0.80848
|  | 0.1725 | 0.98982 | 0.80887
Item # | NEL03A(/M) | 0.20111 | 0.98871 | 0.8089
DISCLAIMER: The data presented here are typical. Slight variations in performance data will occur from lot to lot. Please contact Technical Support with any questions regarding the use or reliability of this data. |  | 0.23 | 0.97857 | 0.80846
|  | 0.25861 | 0.98773 | 0.80877
|  | 0.28722 | 0.99163 | 0.80894
|  | 0.31611 | 0.98013 | 0.8086
|  | 0.345 | 0.98518 | 0.80869
|  | 0.37361 | 0.99354 | 0.80899
|  | 0.40222 | 0.98273 | 0.80867
Additional Information: |  | 0.43083 | 0.9828 | 0.80857
|  | 0.45944 | 0.99407 | 0.80889
|  | 0.48806 | 0.98631 | 0.80884
|  | 0.51667 | 0.98123 | 0.80859
|  | 0.54528 | 0.99327 | 0.80889
|  | 0.57417 | 0.99014 | 0.80883
|  | 0.60278 | 0.9809 | 0.8086
|  | 0.63139 | 0.99127 | 0.80885
|  | 0.66 | 0.99343 | 0.80883
|  | 0.68889 | 0.98201 | 0.80865
|  | 0.7175 | 0.98872 | 0.80873
|  | 0.74611 | 0.99566 | 0.80875
|  | 0.77472 | 0.98438 | 0.8087
|  | 0.80333 | 0.98609 | 0.80871
|  | 0.83194 | 0.99643 | 0.80868
|  | 0.86083 | 0.98744 | 0.80868
|  | 0.88944 | 0.98398 | 0.80864
|  | 0.91806 | 0.99568 | 0.80868
|  | 0.94667 | 0.99086 | 0.80864
|  | 0.97528 | 0.98293 | 0.80859
|  | 1.0042 | 0.99376 | 0.80853
|  | 1.0328 | 0.99387 | 0.80855
|  | 1.0614 | 0.98315 | 0.80853
|  | 1.09 | 0.99103 | 0.80848
|  | 1.1186 | 0.99609 | 0.80851
|  | 1.1478 | 0.98457 | 0.8085
|  | 1.1767 | 0.98833 | 0.80849
|  | 1.2053 | 0.9973 | 0.80854
|  | 1.2342 | 0.9871 | 0.80846
|  | 1.2628 | 0.98592 | 0.80844
|  | 1.2914 | 0.99717 | 0.80855
|  | 1.3203 | 0.99027 | 0.80839
|  | 1.3489 | 0.98434 | 0.80842
|  | 1.3775 | 0.9958 | 0.80854
|  | 1.4061 | 0.99358 | 0.80845
|  | 1.4347 | 0.98387 | 0.80833
|  | 1.4636 | 0.99335 | 0.80839
|  | 1.4922 | 0.9963 | 0.80856
|  | 1.5208 | 0.98488 | 0.80835
|  | 1.5494 | 0.99054 | 0.80839
|  | 1.5783 | 0.99794 | 0.80862
|  | 1.6069 | 0.987 | 0.80838
|  | 1.6358 | 0.98788 | 0.80831
|  | 1.6644 | 0.99835 | 0.80865
|  | 1.6933 | 0.98993 | 0.80836
|  | 1.7219 | 0.98573 | 0.80837
|  | 1.7506 | 0.99736 | 0.80863
|  | 1.7794 | 0.99309 | 0.80845
|  | 1.8081 | 0.98476 | 0.80836
|  | 1.8367 | 0.99526 | 0.80853
|  | 1.8653 | 0.99601 | 0.80859
|  | 1.8942 | 0.98517 | 0.80831
|  | 1.9228 | 0.99247 | 0.8085
|  | 1.9514 | 0.9981 | 0.80868
|  | 1.98 | 0.98678 | 0.80834
|  | 2.0086 | 0.98957 | 0.80831
|  | 2.0375 | 0.99887 | 0.80871
|  | 2.0661 | 0.98928 | 0.80832
|  | 2.0947 | 0.98711 | 0.80836
|  | 2.1236 | 0.99844 | 0.80868
|  | 2.1528 | 0.99208 | 0.80845
|  | 2.1814 | 0.9856 | 0.80834
|  | 2.2103 | 0.99676 | 0.80861
|  | 2.2389 | 0.99505 | 0.80855
|  | 2.2678 | 0.9852 | 0.80831
|  | 2.2964 | 0.99425 | 0.80852
|  | 2.325 | 0.99743 | 0.8086
|  | 2.3539 | 0.98611 | 0.80829
|  | 2.3825 | 0.99127 | 0.80846
|  | 2.4111 | 0.99879 | 0.80874
|  | 2.44 | 0.98806 | 0.80845
|  | 2.4686 | 0.98849 | 0.80838
|  | 2.4972 | 0.99894 | 0.80863
|  | 2.5261 | 0.99091 | 0.80846
|  | 2.5547 | 0.98648 | 0.8083
|  | 2.5833 | 0.99781 | 0.80857
|  | 2.6119 | 0.99406 | 0.80844
|  | 2.6406 | 0.98563 | 0.80839
|  | 2.6694 | 0.99561 | 0.80855
|  | 2.6981 | 0.99674 | 0.80866
|  | 2.7267 | 0.98601 | 0.80824
|  | 2.7553 | 0.99278 | 0.80848
|  | 2.7842 | 0.99856 | 0.80882
|  | 2.8128 | 0.98769 | 0.80829
|  | 2.8414 | 0.98989 | 0.8084
|  | 2.8703 | 0.99922 | 0.80894
|  | 2.8989 | 0.99026 | 0.80841
|  | 2.9275 | 0.98756 | 0.80829
|  | 2.9561 | 0.99851 | 0.80893
|  | 2.9847 | 0.99328 | 0.80867
|  | 3.0133 | 0.98618 | 0.80828
|  | 3.0419 | 0.99669 | 0.8089
|  | 3.0706 | 0.99616 | 0.80887
|  | 3.0994 | 0.98611 | 0.80831
|  | 3.1286 | 0.99429 | 0.8088
|  | 3.1572 | 0.9982 | 0.80895
|  | 3.1858 | 0.98719 | 0.8084
|  | 3.2147 | 0.9914 | 0.80868
|  | 3.2433 | 0.99931 | 0.80893
|  | 3.2719 | 0.98931 | 0.80866
|  | 3.3006 | 0.98878 | 0.80859
|  | 3.3294 | 0.99917 | 0.80895
|  | 3.3581 | 0.99225 | 0.80876
|  | 3.3867 | 0.98693 | 0.80857
|  | 3.4153 | 0.99786 | 0.80889
|  | 3.4439 | 0.9952 | 0.80893
|  | 3.4728 | 0.98638 | 0.80862
|  | 3.5014 | 0.99564 | 0.80884
|  | 3.53 | 0.99778 | 0.80878
|  | 3.5586 | 0.98708 | 0.80862
|  | 3.5875 | 0.9929 | 0.80878
|  | 3.6161 | 0.99935 | 0.80868
|  | 3.6447 | 0.98893 | 0.80871
|  | 3.6736 | 0.99017 | 0.80871
|  | 3.7022 | 0.9997 | 0.80865
|  | 3.7308 | 0.99149 | 0.80876
|  | 3.7597 | 0.98804 | 0.80867
|  | 3.7883 | 0.99874 | 0.80861
|  | 3.8169 | 0.99443 | 0.80864
|  | 3.8456 | 0.98686 | 0.80866
|  | 3.8742 | 0.99681 | 0.80857
|  | 3.9031 | 0.99704 | 0.80855
|  | 3.9317 | 0.98698 | 0.80863
|  | 3.9603 | 0.99412 | 0.80853
|  | 3.9892 | 0.99886 | 0.80853
|  | 4.0178 | 0.98823 | 0.80865
|  | 4.0467 | 0.99137 | 0.80853
|  | 4.0753 | 0.9996 | 0.80851
|  | 4.1039 | 0.99043 | 0.80854
|  | 4.1331 | 0.98906 | 0.80852
|  | 4.1617 | 0.99925 | 0.80849
|  | 4.1906 | 0.99296 | 0.80844
|  | 4.2192 | 0.98735 | 0.80848
|  | 4.2481 | 0.99781 | 0.80843
|  | 4.2767 | 0.99564 | 0.80835
|  | 4.3053 | 0.98681 | 0.80841
|  | 4.3342 | 0.9954 | 0.80841
|  | 4.3628 | 0.99777 | 0.80848
|  | 4.3914 | 0.9874 | 0.8084
|  | 4.4203 | 0.99263 | 0.80832
|  | 4.4489 | 0.99911 | 0.80853
|  | 4.4778 | 0.98904 | 0.80831
|  | 4.5064 | 0.98992 | 0.80829
|  | 4.535 | 0.99927 | 0.80853
|  | 4.5636 | 0.99158 | 0.80829
|  | 4.5925 | 0.98777 | 0.80831
|  | 4.6211 | 0.99827 | 0.80857
|  | 4.6497 | 0.99419 | 0.80834
|  | 4.6786 | 0.98663 | 0.80829
|  | 4.7072 | 0.99619 | 0.80841
|  | 4.7361 | 0.99665 | 0.80847
|  | 4.7647 | 0.98674 | 0.80826
|  | 4.7933 | 0.9936 | 0.80821
|  | 4.8219 | 0.99839 | 0.80855
|  | 4.8508 | 0.98801 | 0.80822
|  | 4.8794 | 0.9908 | 0.80819
|  | 4.9081 | 0.99907 | 0.8087
|  | 4.9369 | 0.99019 | 0.80816
|  | 4.9656 | 0.98836 | 0.80812
|  | 4.9942 | 0.99852 | 0.80865
|  | 5.0228 | 0.99286 | 0.80828
|  | 5.0514 | 0.98679 | 0.80812
|  | 5.0803 | 0.99683 | 0.80854
|  | 5.1089 | 0.99542 | 0.8085
|  | 5.1378 | 0.98639 | 0.80809
|  | 5.1667 | 0.9946 | 0.80843
|  | 5.1953 | 0.99745 | 0.80864
|  | 5.2239 | 0.98715 | 0.80805
|  | 5.2528 | 0.99182 | 0.8083
|  | 5.2814 | 0.99856 | 0.80872
|  | 5.3103 | 0.98891 | 0.80812
|  | 5.3389 | 0.98916 | 0.80813
|  | 5.3675 | 0.99858 | 0.80876
|  | 5.3961 | 0.99146 | 0.80828
|  | 5.4247 | 0.98715 | 0.8081
|  | 5.4536 | 0.99749 | 0.80861
|  | 5.4822 | 0.99413 | 0.80846
|  | 5.5111 | 0.98628 | 0.80806
|  | 5.5397 | 0.99543 | 0.80852
|  | 5.5683 | 0.99649 | 0.80863
|  | 5.5969 | 0.98653 | 0.80809
|  | 5.6258 | 0.99274 | 0.8084
|  | 5.6544 | 0.99804 | 0.80875
|  | 5.6831 | 0.98798 | 0.80816
|  | 5.7117 | 0.98999 | 0.80825
|  | 5.7403 | 0.99857 | 0.80874
|  | 5.7689 | 0.99033 | 0.8083
|  | 5.7978 | 0.98775 | 0.8081
|  | 5.8264 | 0.9979 | 0.80871
|  | 5.8553 | 0.99301 | 0.80846
|  | 5.8839 | 0.98649 | 0.80806
|  | 5.9125 | 0.99621 | 0.80861
|  | 5.9411 | 0.9955 | 0.80862
|  | 5.97 | 0.98624 | 0.80805
|  | 5.9986 | 0.99376 | 0.80851
|  | 6.0272 | 0.99744 | 0.80873
|  | 6.0561 | 0.98729 | 0.80813
|  | 6.0847 | 0.99103 | 0.80839
|  | 6.1133 | 0.99831 | 0.80871
|  | 6.1436 | 0.98865 | 0.8082
|  | 6.1725 | 0.98912 | 0.80822
|  | 6.2011 | 0.99841 | 0.80878
|  | 6.2297 | 0.99113 | 0.80838
|  | 6.2586 | 0.98711 | 0.80813
|  | 6.2872 | 0.99736 | 0.80873
|  | 6.3158 | 0.99375 | 0.8086
|  | 6.3447 | 0.98618 | 0.8081
|  | 6.3733 | 0.99542 | 0.80865
|  | 6.4019 | 0.99625 | 0.80874
|  | 6.4308 | 0.98647 | 0.80817
|  | 6.4594 | 0.99286 | 0.80861
|  | 6.4881 | 0.99788 | 0.80876
|  | 6.5169 | 0.98782 | 0.80829
|  | 6.5456 | 0.99019 | 0.80838
|  | 6.5744 | 0.99846 | 0.80874
|  | 6.6031 | 0.99006 | 0.8084
|  | 6.6317 | 0.98787 | 0.80825
|  | 6.6603 | 0.99789 | 0.80876
|  | 6.6889 | 0.99273 | 0.80857
|  | 6.7178 | 0.98654 | 0.80818
|  | 6.7464 | 0.99636 | 0.80871
|  | 6.775 | 0.9953 | 0.80871
|  | 6.8039 | 0.98626 | 0.80817
|  | 6.8325 | 0.99392 | 0.80859
|  | 6.8611 | 0.99734 | 0.80873
|  | 6.8897 | 0.98717 | 0.80821
|  | 6.9183 | 0.99119 | 0.8085
|  | 6.9472 | 0.99836 | 0.80874
|  | 6.9758 | 0.98908 | 0.80837
|  | 7.0044 | 0.98875 | 0.80837
|  | 7.0333 | 0.99835 | 0.8088
|  | 7.0619 | 0.99165 | 0.80855
|  | 7.0908 | 0.98698 | 0.80825
|  | 7.1194 | 0.99719 | 0.80875
|  | 7.1483 | 0.99415 | 0.80864
|  | 7.1772 | 0.9862 | 0.80819
|  | 7.2058 | 0.99505 | 0.80866
|  | 7.2344 | 0.99648 | 0.80876
|  | 7.2631 | 0.9866 | 0.80821
|  | 7.2917 | 0.99251 | 0.80857
|  | 7.3206 | 0.99805 | 0.80879
|  | 7.3492 | 0.98818 | 0.8083
|  | 7.3778 | 0.98985 | 0.80844
|  | 7.4064 | 0.99854 | 0.80874
|  | 7.4353 | 0.99052 | 0.80855
|  | 7.4639 | 0.98767 | 0.80828
|  | 7.4925 | 0.9978 | 0.80877
|  | 7.5211 | 0.99327 | 0.80861
|  | 7.5497 | 0.98642 | 0.80822
|  | 7.5786 | 0.99594 | 0.80871
|  | 7.6072 | 0.9958 | 0.8087
|  | 7.6358 | 0.98641 | 0.80826
|  | 7.6644 | 0.99355 | 0.80865
|  | 7.6933 | 0.99776 | 0.80877
|  | 7.7219 | 0.98763 | 0.80836
|  | 7.7506 | 0.99082 | 0.80854
|  | 7.7792 | 0.99854 | 0.80875
|  | 7.8081 | 0.98968 | 0.80847
|  | 7.8367 | 0.98847 | 0.8084
|  | 7.8653 | 0.99822 | 0.80874
|  | 7.8939 | 0.99228 | 0.80861
|  | 7.9228 | 0.98683 | 0.80832
|  | 7.9514 | 0.9968 | 0.80869
|  | 7.98 | 0.99491 | 0.80872
|  | 8.0086 | 0.98635 | 0.80821
|  | 8.0375 | 0.99455 | 0.8087
|  | 8.0661 | 0.99708 | 0.80875
|  | 8.0947 | 0.98703 | 0.80831
|  | 8.1236 | 0.99196 | 0.80857
|  | 8.1525 | 0.99834 | 0.80879
|  | 8.1811 | 0.98873 | 0.80841
|  | 8.2097 | 0.98941 | 0.80851
|  | 8.2386 | 0.99855 | 0.80878
|  | 8.2672 | 0.99119 | 0.80859
|  | 8.2958 | 0.98744 | 0.80833
|  | 8.3244 | 0.99752 | 0.80874
|  | 8.3531 | 0.99393 | 0.80872
|  | 8.3817 | 0.98643 | 0.80828
|  | 8.4103 | 0.99553 | 0.80874
|  | 8.4392 | 0.99632 | 0.80871
|  | 8.4678 | 0.98666 | 0.80831
|  | 8.4967 | 0.99304 | 0.80865
|  | 8.5253 | 0.998 | 0.80872
|  | 8.5542 | 0.98802 | 0.80843
|  | 8.5828 | 0.99044 | 0.80856
|  | 8.6114 | 0.99863 | 0.80868
|  | 8.64 | 0.99032 | 0.80856
|  | 8.6686 | 0.9882 | 0.8084
|  | 8.6975 | 0.99808 | 0.80871
|  | 8.7261 | 0.99287 | 0.80868
|  | 8.755 | 0.98677 | 0.8083
|  | 8.7836 | 0.99653 | 0.80872
|  | 8.8125 | 0.99535 | 0.80872
|  | 8.8411 | 0.98647 | 0.80825
|  | 8.8697 | 0.99427 | 0.8087
|  | 8.8986 | 0.99739 | 0.80877
|  | 8.9272 | 0.98737 | 0.80831
|  | 8.9558 | 0.99166 | 0.80859
|  | 8.9847 | 0.9985 | 0.80874
|  | 9.0133 | 0.98924 | 0.80846
|  | 9.0422 | 0.98922 | 0.80844
|  | 9.0708 | 0.99854 | 0.80873
|  | 9.0997 | 0.99164 | 0.80861
|  | 9.1286 | 0.9875 | 0.80838
|  | 9.1575 | 0.99756 | 0.80875
|  | 9.1861 | 0.99418 | 0.80876
|  | 9.2147 | 0.98661 | 0.8083
|  | 9.2433 | 0.99555 | 0.80875
|  | 9.2722 | 0.99656 | 0.80877
|  | 9.3008 | 0.98688 | 0.80833
|  | 9.3297 | 0.99303 | 0.80865
|  | 9.3583 | 0.99815 | 0.80875
|  | 9.3869 | 0.98821 | 0.80841
|  | 9.4158 | 0.99039 | 0.80855
|  | 9.4444 | 0.99871 | 0.80873
|  | 9.4731 | 0.99044 | 0.80858
|  | 9.5019 | 0.98829 | 0.80838
|  | 9.5306 | 0.99824 | 0.80873
|  | 9.5594 | 0.99309 | 0.80867
|  | 9.5881 | 0.98698 | 0.80835
|  | 9.6167 | 0.99665 | 0.80878
|  | 9.6456 | 0.99564 | 0.80875
|  | 9.6742 | 0.98666 | 0.80832
|  | 9.7028 | 0.99422 | 0.80872
|  | 9.7314 | 0.99764 | 0.80875
|  | 9.76 | 0.98762 | 0.80837
|  | 9.7886 | 0.99157 | 0.8086
|  | 9.8175 | 0.99872 | 0.80872
|  | 9.8461 | 0.98961 | 0.80854
|  | 9.8747 | 0.98914 | 0.80852
|  | 9.9033 | 0.99862 | 0.80876
|  | 9.9319 | 0.99218 | 0.80863
|  | 9.9606 | 0.98739 | 0.8084
|  | 9.9894 | 0.99731 | 0.80875
|  | 10.0181 | 0.99479 | 0.80869
|  | 10.0469 | 0.98667 | 0.80833
|  | 10.0756 | 0.99525 | 0.80869
|  | 10.1044 | 0.997 | 0.80877
|  | 10.1333 | 0.98715 | 0.80843
|  | 10.1619 | 0.99285 | 0.80863
|  | 10.1908 | 0.99848 | 0.80874
|  | 10.2194 | 0.98873 | 0.80849
|  | 10.2481 | 0.99027 | 0.80855
|  | 10.2767 | 0.99882 | 0.80873
|  | 10.3056 | 0.99104 | 0.80855
|  | 10.3342 | 0.98812 | 0.8084
|  | 10.3628 | 0.99811 | 0.80868
|  | 10.3914 | 0.99369 | 0.80869
|  | 10.4203 | 0.98699 | 0.80836
|  | 10.4489 | 0.9964 | 0.80875
|  | 10.4778 | 0.99613 | 0.80876
|  | 10.5064 | 0.98698 | 0.80835
|  | 10.5353 | 0.99403 | 0.80868
|  | 10.5639 | 0.99791 | 0.80877
|  | 10.5925 | 0.98797 | 0.80837
|  | 10.6214 | 0.99144 | 0.80856
|  | 10.65 | 0.9989 | 0.8087
|  | 10.6786 | 0.99003 | 0.8085
|  | 10.7072 | 0.98908 | 0.80846
|  | 10.7361 | 0.99866 | 0.80874
|  | 10.7647 | 0.99257 | 0.80868
|  | 10.7936 | 0.98753 | 0.80838
|  | 10.8222 | 0.99736 | 0.80876
|  | 10.8511 | 0.9952 | 0.80869
|  | 10.8797 | 0.98696 | 0.80831
|  | 10.9083 | 0.99522 | 0.80876
|  | 10.9369 | 0.99732 | 0.80874
|  | 10.9658 | 0.98747 | 0.80834
|  | 10.9944 | 0.99261 | 0.8086
|  | 11.0231 | 0.9987 | 0.8087
|  | 11.0519 | 0.98912 | 0.80843
|  | 11.0806 | 0.9902 | 0.8085
|  | 11.1094 | 0.99898 | 0.80881
|  | 11.1383 | 0.99142 | 0.80861
|  | 11.1672 | 0.98827 | 0.80834
|  | 11.1958 | 0.99813 | 0.80871
|  | 11.2244 | 0.99407 | 0.8087
|  | 11.2531 | 0.9871 | 0.80831
|  | 11.2817 | 0.99637 | 0.80875
|  | 11.3103 | 0.99658 | 0.80878
|  | 11.3389 | 0.98724 | 0.80834
|  | 11.3675 | 0.99386 | 0.80869
|  | 11.3964 | 0.99837 | 0.80873
|  | 11.425 | 0.9885 | 0.80842
|  | 11.4536 | 0.99128 | 0.80858
|  | 11.4825 | 0.99909 | 0.80873
|  | 11.5111 | 0.99056 | 0.80852
|  | 11.5397 | 0.98897 | 0.80846
|  | 11.5686 | 0.99861 | 0.80869
|  | 11.5972 | 0.99306 | 0.8087
|  | 11.6261 | 0.98753 | 0.80838
|  | 11.6547 | 0.99727 | 0.8087
|  | 11.6833 | 0.99573 | 0.80873
|  | 11.7122 | 0.98717 | 0.80832
|  | 11.7408 | 0.99506 | 0.80875
|  | 11.7694 | 0.99781 | 0.80877
|  | 11.7981 | 0.98795 | 0.80839
|  | 11.8267 | 0.99238 | 0.8086
|  | 11.8556 | 0.99897 | 0.80872
|  | 11.8842 | 0.98966 | 0.80851
|  | 11.9131 | 0.99006 | 0.80849
|  | 11.9417 | 0.9991 | 0.80877
|  | 11.9706 | 0.99211 | 0.80863
|  | 11.9992 | 0.98824 | 0.80835
|  | 12.0278 | 0.99816 | 0.80874
|  | 12.0567 | 0.99474 | 0.80873
|  | 12.0853 | 0.98726 | 0.80833
|  | 12.1139 | 0.99624 | 0.80876
|  | 12.1431 | 0.99685 | 0.80871
|  | 12.1719 | 0.98744 | 0.80838
|  | 12.2006 | 0.99394 | 0.80871
|  | 12.2294 | 0.99852 | 0.80876
|  | 12.2581 | 0.98873 | 0.80846
|  | 12.2867 | 0.99129 | 0.80861
|  | 12.3153 | 0.99924 | 0.80871
|  | 12.3442 | 0.99087 | 0.80859
|  | 12.3728 | 0.98908 | 0.80843
|  | 12.4014 | 0.99877 | 0.80874
|  | 12.43 | 0.99345 | 0.80869
|  | 12.4589 | 0.98759 | 0.80833
|  | 12.4875 | 0.99726 | 0.80874
|  | 12.5164 | 0.99587 | 0.80875
|  | 12.545 | 0.98727 | 0.80833
|  | 12.5739 | 0.99511 | 0.80874
|  | 12.6028 | 0.9979 | 0.80881
|  | 12.6314 | 0.98807 | 0.80838
|  | 12.66 | 0.99254 | 0.80865
|  | 12.6889 | 0.99911 | 0.80874
|  | 12.7175 | 0.98982 | 0.80848
|  | 12.7461 | 0.99008 | 0.8085
|  | 12.775 | 0.99915 | 0.80877
|  | 12.8036 | 0.9922 | 0.80863
|  | 12.8325 | 0.98828 | 0.8084
|  | 12.8611 | 0.99817 | 0.8088
|  | 12.89 | 0.99483 | 0.80871
|  | 12.9186 | 0.98741 | 0.80835
|  | 12.9472 | 0.99636 | 0.80876
|  | 12.9761 | 0.99718 | 0.8088
|  | 13.0047 | 0.98762 | 0.80832
|  | 13.0336 | 0.9939 | 0.80865
|  | 13.0622 | 0.99874 | 0.80877
|  | 13.0911 | 0.98892 | 0.80842
|  | 13.1197 | 0.99139 | 0.80853
|  | 13.1489 | 0.99938 | 0.80876
|  | 13.1775 | 0.99096 | 0.8085
|  | 13.2064 | 0.98929 | 0.80844
|  | 13.235 | 0.99898 | 0.8088
|  | 13.2636 | 0.99352 | 0.80866
|  | 13.2925 | 0.98786 | 0.80832
|  | 13.3211 | 0.99753 | 0.80876
|  | 13.35 | 0.99603 | 0.80877
|  | 13.3786 | 0.98746 | 0.8083
|  | 13.4075 | 0.99535 | 0.80872
|  | 13.4361 | 0.99807 | 0.80879
|  | 13.4647 | 0.98825 | 0.80832
|  | 13.4936 | 0.99281 | 0.80858
|  | 13.5222 | 0.99931 | 0.80879
|  | 13.5508 | 0.99008 | 0.80843
|  | 13.5797 | 0.99041 | 0.80848
|  | 13.6083 | 0.99947 | 0.80874
|  | 13.6372 | 0.99239 | 0.80858
|  | 13.6658 | 0.98854 | 0.80837
|  | 13.6944 | 0.99849 | 0.8088
|  | 13.7233 | 0.99502 | 0.80871
|  | 13.7519 | 0.98769 | 0.8083
|  | 13.7808 | 0.99665 | 0.80876
|  | 13.8094 | 0.99739 | 0.80877
|  | 13.8381 | 0.98793 | 0.80833
|  | 13.8669 | 0.99418 | 0.80862
|  | 13.8956 | 0.99905 | 0.80878
|  | 13.9244 | 0.98925 | 0.80839
|  | 13.9531 | 0.99166 | 0.80851
|  | 13.9819 | 0.99965 | 0.8088
|  | 14.0106 | 0.99134 | 0.80852
|  | 14.0392 | 0.98946 | 0.80839
|  | 14.0681 | 0.9992 | 0.80877
|  | 14.0967 | 0.99393 | 0.80861
|  | 14.1256 | 0.98833 | 0.80827
|  | 14.1556 | 0.99813 | 0.80879
|  | 14.1842 | 0.99596 | 0.80873
|  | 14.2131 | 0.98778 | 0.80823
|  | 14.2417 | 0.99602 | 0.80871
|  | 14.2703 | 0.99813 | 0.8088
|  | 14.2992 | 0.98831 | 0.80824
|  | 14.3278 | 0.99345 | 0.80854
|  | 14.3567 | 0.99943 | 0.8088
|  | 14.3853 | 0.98984 | 0.80835
|  | 14.4142 | 0.99097 | 0.80838
|  | 14.4428 | 0.99971 | 0.8088
|  | 14.4717 | 0.9922 | 0.80846
|  | 14.5003 | 0.989 | 0.80826
|  | 14.5289 | 0.99891 | 0.8088
|  | 14.5575 | 0.99486 | 0.80863
|  | 14.5864 | 0.98795 | 0.80824
|  | 14.615 | 0.99717 | 0.80872
|  | 14.6439 | 0.99728 | 0.8087
|  | 14.6725 | 0.98794 | 0.80817
|  | 14.7011 | 0.99471 | 0.80854
|  | 14.73 | 0.99912 | 0.80875
|  | 14.7586 | 0.98923 | 0.8083
|  | 14.7872 | 0.99219 | 0.80848
|  | 14.8158 | 1 | 0.80884
|  | 14.8447 | 0.99138 | 0.8084
|  | 14.8733 | 0.96338 | 0.80837

## Transmission
|  | Wavelength (nm) | Transmission (%)
|  | 500 | 53.6255
|  | 501 | 54.0713
|  | 502 | 54.6295
|  | 503 | 55.1741
Noise Eater Transmission |  | 504 | 55.773
|  | 505 | 56.3332
|  | 506 | 56.9549
Item # | NEL03A(/M) | 507 | 57.4517
DISCLAIMER: The data presented here are typical. Slight variations in performance data will occur from lot to lot. Please contact Technical Support with any questions regarding the use or reliability of this data. |  | 508 | 57.6319
|  | 509 | 57.7839
|  | 510 | 58.1432
|  | 511 | 58.489
|  | 512 | 58.8482
|  | 513 | 58.8555
|  | 514 | 59.0336
Additional Information: |  | 515 | 59.647
|  | 516 | 60.4177
|  | 517 | 60.9457
|  | 518 | 61.2122
|  | 519 | 61.5997
|  | 520 | 61.8402
|  | 521 | 62.3255
|  | 522 | 62.5026
|  | 523 | 62.8639
|  | 524 | 63.0181
|  | 525 | 63.4461
|  | 526 | 63.9158
|  | 527 | 63.998
|  | 528 | 64.1522
|  | 529 | 64.6052
|  | 530 | 65.1051
|  | 531 | 65.3977
|  | 532 | 65.5789
|  | 533 | 65.9476
|  | 534 | 66.0517
|  | 535 | 66.1048
|  | 536 | 66.1694
|  | 537 | 66.435
|  | 538 | 66.8151
|  | 539 | 67.1848
|  | 540 | 67.5336
|  | 541 | 67.7711
|  | 542 | 67.7825
|  | 543 | 68.1835
|  | 544 | 68.2324
|  | 545 | 68.5719
|  | 546 | 68.4824
|  | 547 | 68.9427
|  | 548 | 69.2426
|  | 549 | 69.6456
|  | 550 | 69.6206
|  | 551 | 69.7394
|  | 552 | 69.7529
|  | 553 | 69.9206
|  | 554 | 70.1965
|  | 555 | 70.5902
|  | 556 | 70.9745
|  | 557 | 70.9578
|  | 558 | 71.2577
|  | 559 | 71.0849
|  | 560 | 71.3234
|  | 561 | 71.214
|  | 562 | 71.5774
|  | 563 | 71.7566
|  | 564 | 72.0315
|  | 565 | 72.4012
|  | 566 | 72.772
|  | 567 | 73.1073
|  | 568 | 73.779
|  | 569 | 73.7509
|  | 570 | 73.6384
|  | 571 | 73.1739
|  | 572 | 73.3343
|  | 573 | 73.3093
|  | 574 | 73.2281
|  | 575 | 73.4405
|  | 576 | 73.6301
|  | 577 | 73.8811
|  | 578 | 74.1362
|  | 579 | 74.4132
|  | 580 | 74.5798
|  | 581 | 74.4663
|  | 582 | 74.6434
|  | 583 | 74.8631
|  | 584 | 75.2193
|  | 585 | 75.4463
|  | 586 | 75.7827
|  | 587 | 75.6098
|  | 588 | 75.7087
|  | 589 | 75.5609
|  | 590 | 76.016
|  | 591 | 75.991
|  | 592 | 76.0305
|  | 593 | 76.2544
|  | 594 | 76.5887
|  | 595 | 76.9709
|  | 596 | 76.9959
|  | 597 | 76.5846
|  | 598 | 76.4002
|  | 599 | 76.6668
|  | 600 | 76.9418
|  | 601 | 76.9647
|  | 602 | 76.4388
|  | 603 | 76.7908
|  | 604 | 77.4031
|  | 605 | 77.6708
|  | 606 | 77.7166
|  | 607 | 77.8093
|  | 608 | 77.9322
|  | 609 | 77.8624
|  | 610 | 77.651
|  | 611 | 77.8582
|  | 612 | 78.1769
|  | 613 | 78.0717
|  | 614 | 78.1384
|  | 615 | 78.2342
|  | 616 | 78.2165
|  | 617 | 78.0363
|  | 618 | 78.1186
|  | 619 | 78.204
|  | 620 | 78.3852
|  | 621 | 78.7268
|  | 622 | 79.2766
|  | 623 | 79.3849
|  | 624 | 79.2579
|  | 625 | 79.4589
|  | 626 | 79.6599
|  | 627 | 79.9619
|  | 628 | 79.5755
|  | 629 | 79.6151
|  | 630 | 79.2527
|  | 631 | 79.4662
|  | 632 | 79.8025
|  | 633 | 79.9629
|  | 634 | 80.4357
|  | 635 | 80.1379
|  | 636 | 80.0369
|  | 637 | 79.4735
|  | 638 | 79.7942
|  | 639 | 79.7369
|  | 640 | 79.738
|  | 641 | 80.0525
|  | 642 | 80.5732
|  | 643 | 81.2001
|  | 644 | 81.2709
|  | 645 | 81.1397
|  | 646 | 80.9512
|  | 647 | 80.7294
|  | 648 | 80.6628
|  | 649 | 80.5774
|  | 650 | 80.3941
|  | 651 | 80.5232
|  | 652 | 80.6523
|  | 653 | 80.7815
|  | 654 | 80.9106
|  | 655 | 81.0397
|  | 656 | 81.1699
|  | 657 | 81.0075
|  | 658 | 81.2303
|  | 659 | 81.2282
|  | 660 | 81.2407
|  | 661 | 81.5417
|  | 662 | 81.4594
|  | 663 | 81.7989
|  | 664 | 81.9229
|  | 665 | 81.9989
|  | 666 | 81.9499
|  | 667 | 81.8593
|  | 668 | 81.8625
|  | 669 | 81.9739
|  | 670 | 82.2728
|  | 671 | 82.4279
|  | 672 | 82.0968
|  | 673 | 82.1655
|  | 674 | 82.4758
|  | 675 | 82.7102
|  | 676 | 82.6008
|  | 677 | 82.0405
|  | 678 | 82.0124
|  | 679 | 81.6729
|  | 680 | 82.0093
|  | 681 | 82.2655
|  | 682 | 82.8112
|  | 683 | 82.8601
|  | 684 | 82.9986
|  | 685 | 82.7518
|  | 686 | 82.9341
|  | 687 | 82.7612
|  | 688 | 82.7924
|  | 689 | 82.7539
|  | 690 | 82.9976
|  | 691 | 82.9185
|  | 692 | 82.8674
|  | 693 | 83.0361
|  | 694 | 83.4621
|  | 695 | 83.1017
|  | 696 | 82.7789
|  | 697 | 82.6862
|  | 698 | 82.9216
|  | 699 | 83.1799
|  | 700 | 83.2528
|  | 701 | 83.1653
|  | 702 | 82.8674
|  | 703 | 82.982
|  | 704 | 83.3131
|  | 705 | 83.5537
|  | 706 | 83.6235
|  | 707 | 83.8682
|  | 708 | 83.9495
|  | 709 | 83.9463
|  | 710 | 83.8141
|  | 711 | 84.039
|  | 712 | 83.812
|  | 713 | 83.8276
|  | 714 | 83.4871
|  | 715 | 83.7787
|  | 716 | 83.8026
|  | 717 | 83.5121
|  | 718 | 83.2403
|  | 719 | 83.3579
|  | 720 | 83.6651
|  | 721 | 83.7287
|  | 722 | 83.7183
|  | 723 | 84.1734
|  | 724 | 84.416
|  | 725 | 84.3702
|  | 726 | 84.341
|  | 727 | 84.4275
|  | 728 | 84.3108
|  | 729 | 84.2775
|  | 730 | 84.0328
|  | 731 | 84.3879
|  | 732 | 84.2817
|  | 733 | 84.1036
|  | 734 | 83.8953
|  | 735 | 83.7328
|  | 736 | 84.1505
|  | 737 | 84.5243
|  | 738 | 84.7763
|  | 739 | 85.1898
|  | 740 | 84.9409
|  | 741 | 84.8211
|  | 742 | 84.7128
|  | 743 | 84.6451
|  | 744 | 84.5649
|  | 745 | 84.3577
|  | 746 | 84.4483
|  | 747 | 84.641
|  | 748 | 84.4691
|  | 749 | 84.7003
|  | 750 | 84.6889
|  | 751 | 84.9773
|  | 752 | 84.7357
|  | 753 | 84.8222
|  | 754 | 84.8284
|  | 755 | 84.8586
|  | 756 | 84.8149
|  | 757 | 84.4545
|  | 758 | 84.7649
|  | 759 | 84.567
|  | 760 | 84.9461
|  | 761 | 84.692
|  | 762 | 85.2325
|  | 763 | 85.2325
|  | 764 | 85.4272
|  | 765 | 85.1085
|  | 766 | 85.3908
|  | 767 | 85.2554
|  | 768 | 85.522
|  | 769 | 85.6001
|  | 770 | 85.5387
|  | 771 | 85.3252
|  | 772 | 84.8211
|  | 773 | 84.8617
|  | 774 | 85.0221
|  | 775 | 85.4199
|  | 776 | 85.4449
|  | 777 | 85.0731
|  | 778 | 84.9607
|  | 779 | 85.0773
|  | 780 | 85.5543
|  | 781 | 85.3658
|  | 782 | 85.4855
|  | 783 | 85.6251
|  | 784 | 85.9188
|  | 785 | 85.9948
|  | 786 | 85.6147
|  | 787 | 85.8407
|  | 788 | 85.7178
|  | 789 | 85.7376
|  | 790 | 85.6938
|  | 791 | 85.7646
|  | 792 | 85.0763
|  | 796 | 84.9221
|  | 800 | 85.0148
|  | 804 | 85.9865
|  | 808 | 86.0354
|  | 812 | 86.0844
|  | 816 | 86.1312
|  | 820 | 86.202
|  | 824 | 86.2333
|  | 828 | 86.2885
|  | 832 | 86.3478
|  | 836 | 86.4186
|  | 840 | 86.477
|  | 844 | 86.5374
|  | 848 | 86.6134
|  | 852 | 86.6551
|  | 856 | 86.6842
|  | 860 | 86.7269
|  | 864 | 86.7883
|  | 868 | 86.8248
|  | 872 | 86.8508
|  | 876 | 86.9227
|  | 880 | 86.9821
|  | 884 | 87.0289
|  | 888 | 87.0216
|  | 892 | 87.0508
|  | 896 | 87.0904
|  | 900 | 87.1903
|  | 904 | 87.2799
|  | 908 | 87.3924
|  | 912 | 87.3924
|  | 916 | 87.4174
|  | 920 | 87.4684
|  | 924 | 87.5528
|  | 928 | 87.6246
|  | 932 | 87.5954
|  | 936 | 87.4767
|  | 940 | 87.3205
|  | 944 | 87.2924
|  | 948 | 87.5007
|  | 952 | 87.7621
|  | 956 | 87.7423
|  | 960 | 87.6215
|  | 964 | 87.6111
|  | 968 | 87.6808
|  | 972 | 87.7006
|  | 976 | 87.7496
|  | 980 | 87.7215
|  | 984 | 87.835
|  | 988 | 87.7767
|  | 992 | 87.8194
|  | 996 | 87.911
|  | 1000 | 87.8152
|  | 1004 | 88.0141
|  | 1008 | 87.8725
|  | 1012 | 88.0318
|  | 1016 | 87.986
|  | 1020 | 87.9464
|  | 1024 | 88.0839
|  | 1028 | 87.9339
|  | 1032 | 88.0807
|  | 1036 | 88.0443
|  | 1040 | 88.0464
|  | 1044 | 88.0526
|  | 1048 | 88.1047
|  | 1052 | 88.2619
|  | 1056 | 88.1682
|  | 1060 | 88.0151
|  | 1064 | 88.1984
|  | 1068 | 88.4369
|  | 1072 | 88.4286
|  | 1076 | 88.1745
|  | 1080 | 88.0495
|  | 1084 | 88.2182
|  | 1088 | 88.5202
|  | 1092 | 88.7702
|  | 1096 | 88.8337
|  | 1100 | 88.7712
|  | 1104 | 88.6733
|  | 1108 | 88.5515
|  | 1112 | 88.4702
|  | 1116 | 88.338
|  | 1120 | 88.2911
|  | 1124 | 88.2713
|  | 1128 | 88.3526
|  | 1132 | 88.4681
|  | 1136 | 88.6254
|  | 1140 | 88.8379
|  | 1144 | 89.0149
|  | 1148 | 89.1045
|  | 1152 | 88.9753
|  | 1156 | 88.7316
|  | 1160 | 88.4359
|  | 1164 | 88.3005
|  | 1168 | 88.3828
|  | 1172 | 88.7795
|  | 1176 | 89.0763
|  | 1180 | 89.0191
|  | 1184 | 88.6483
|  | 1188 | 88.5171
|  | 1192 | 88.792
|  | 1196 | 88.9107
|  | 1200 | 88.7483
|  | 1204 | 88.616
|  | 1208 | 88.8462
|  | 1212 | 88.8993
|  | 1216 | 88.7472
|  | 1220 | 88.7285
|  | 1224 | 88.7795
|  | 1228 | 88.8087
|  | 1232 | 88.8733
|  | 1236 | 88.8045
|  | 1240 | 88.8691
|  | 1244 | 88.8451
|  | 1248 | 88.8108
|  | 1252 | 88.9732
|  | 1256 | 88.8024
|  | 1260 | 88.9347
|  | 1264 | 88.9889
|  | 1268 | 88.8493
|  | 1272 | 89.1263
|  | 1276 | 88.8045
|  | 1280 | 89.1847
|  | 1284 | 88.8014
|  | 1288 | 89.1794
|  | 1292 | 88.8597
|  | 1296 | 89.2617
|  | 1300 | 88.8805
|  | 1304 | 89.2471
|  | 1308 | 88.8712
|  | 1312 | 89.268
|  | 1316 | 88.9201
|  | 1320 | 89.2096
|  | 1324 | 89.0628
|  | 1328 | 89.1367
|  | 1332 | 89.2492
|  | 1336 | 89.0399
|  | 1340 | 89.3315
|  | 1344 | 89.092
|  | 1348 | 89.194
|  | 1352 | 89.2201
|  | 1356 | 89.1607
|  | 1360 | 89.3013
|  | 1364 | 89.369
|  | 1368 | 89.3284
|  | 1372 | 89.32
|  | 1376 | 89.5012
|  | 1380 | 89.5335
|  | 1384 | 89.5533
|  | 1388 | 89.2971
|  | 1392 | 89.3429
|  | 1396 | 89.1857
|  | 1400 | 89.1722
|  | 1404 | 89.4533
|  | 1408 | 89.4013
|  | 1412 | 89.1326
|  | 1416 | 89.2138
|  | 1420 | 89.5221
|  | 1424 | 89.4346
|  | 1428 | 89.1045
|  | 1432 | 89.1971
|  | 1436 | 89.6168
|  | 1440 | 89.6574
|  | 1444 | 89.2534
|  | 1448 | 88.9639
|  | 1452 | 89.244
|  | 1456 | 89.7366
|  | 1460 | 89.8595
|  | 1464 | 89.4804
|  | 1468 | 88.9962
|  | 1472 | 88.9597
|  | 1476 | 89.3867
|  | 1480 | 89.9022
|  | 1484 | 90.0834
|  | 1488 | 89.8168
|  | 1492 | 89.3752
|  | 1496 | 88.9868
|  | 1500 | 88.8785
|  | 1504 | 89.0566
|  | 1508 | 89.47
|  | 1512 | 89.9199
|  | 1516 | 90.2021
|  | 1520 | 90.2354
|  | 1524 | 90.0219
|  | 1528 | 89.6887
|  | 1532 | 89.345
|  | 1536 | 89.0763
|  | 1540 | 88.8941
|  | 1544 | 88.8222
|  | 1548 | 88.8222
|  | 1552 | 88.9003
|  | 1556 | 89.0545
|  | 1560 | 89.1836
|  | 1564 | 89.3148
|  | 1568 | 89.4086
|  | 1572 | 89.5221
|  | 1576 | 89.596
|  | 1580 | 89.5929
|  | 1584 | 89.6095
|  | 1588 | 89.6179
|  | 1592 | 89.6335
|  | 1596 | 89.6054
|  | 1600 | 89.5585
|  | 1604 | 89.4679
|  | 1608 | 89.3606
|  | 1612 | 89.2044
|  | 1616 | 89.0795
|  | 1620 | 88.9399
|  | 1624 | 88.9097
|  | 1628 | 88.9649
|  | 1632 | 89.1242
|  | 1636 | 89.3461
|  | 1640 | 89.6366
|  | 1644 | 89.9876
|  | 1648 | 90.35
|  | 1652 | 90.6426
|  | 1656 | 90.777
|  | 1660 | 90.6426
|  | 1664 | 90.2938
|  | 1668 | 89.8147
|  | 1672 | 89.4544
|  | 1676 | 89.3179
|  | 1680 | 89.4742
|  | 1684 | 89.8178
|  | 1688 | 90.2198
|  | 1692 | 90.5229
|  | 1696 | 90.552
|  | 1700 | 90.3239
|  | 1704 | 89.9949
|  | 1708 | 89.8126
|  | 1712 | 89.8793
|  | 1716 | 90.1073
|  | 1720 | 90.2896
|  | 1724 | 90.3177
|  | 1728 | 90.2438
|  | 1732 | 90.1625
|  | 1736 | 90.1573
|  | 1740 | 90.1719
|  | 1744 | 90.1959
|  | 1748 | 90.1969
|  | 1752 | 90.1677
|  | 1756 | 90.175
|  | 1760 | 90.1709
|  | 1764 | 90.2156
|  | 1768 | 90.2271
|  | 1772 | 90.2292
|  | 1776 | 90.2011
|  | 1780 | 90.1865
|  | 1784 | 90.1907
|  | 1788 | 90.2042
|  | 1792 | 90.2094
|  | 1796 | 90.2031
|  | 1800 | 90.1927
|  | 1804 | 90.1542
|  | 1808 | 90.1136
|  | 1812 | 90.0626
|  | 1816 | 90.0261
|  | 1820 | 89.9751
|  | 1824 | 89.9043
|  | 1828 | 89.9188
|  | 1832 | 90.0313
|  | 1836 | 90.1823
|  | 1840 | 90.2594
|  | 1844 | 90.2625
|  | 1848 | 90.2948
|  | 1852 | 90.3135
|  | 1856 | 90.3156
|  | 1860 | 90.2875
|  | 1864 | 90.2615
|  | 1868 | 90.249
|  | 1872 | 90.2208
|  | 1876 | 90.2459
|  | 1880 | 90.2386
|  | 1884 | 90.274
|  | 1888 | 90.2979
|  | 1892 | 90.3542
|  | 1896 | 90.376
|  | 1950 | 90.3948
|  | 2000 | 90.603
|  | 2050 | 90.4989
|  | 2100 | 90.7072
|  | 2150 | 90.9155
|  | 2200 | 90.8113
|  | 2250 | 90.9155
|  | 2300 | 91.1238
|  | 2350 | 91.4362
|  | 2400 | 91.6445